| Report No | WO01 Τ016024 01 |
| --- | --- |
| Company name | SPIKA TECH, S.L |
| Audit date(s) | 8&15/7/2025 |

Instructions to the Auditee

- For major non-conformities:

- Within 30 days of audit end date, the client shall submit a response on actions taken to resolve the non-conformities. The response shall include root cause analysis and a plan for implementation of corrections and corrective actions.

- Evidence of effectively implemented corrections and corrective actions shall be provided within 90 days.

- If more days are required to resolve the non-conformities, an extension may be requested, in writing. The request must be accompanied by a proper justification and a schedule for resolving the remaining issues.

- For minor non-conformities:

- Within 30 days of audit end date, the client shall submit a response on actions taken to resolve the non-conformities. The response shall include root cause analysis and a plan for implementation of corrections and corrective actions.

- For observations:

- A written response to observations is encouraged but not mandatory.

- All findings identified are recorded in this form, even if they are corrected during the audit.

- If additional time is required for completion of the corrective actions, an extension period may be granted by HTCert after justified request of the manufacturer.

- At the next audit documented evidence of implementation and effectiveness of the corrective actions have to be presented.

- HTCert reserves the right to conduct follow-up audits to determine that nonconformities have been resolved satisfactorily.

- Please use this form to respond to HTCert with your corrective actions.

This is a confidential report. Its ownership remains with HTCert.

Reproduction is allowed only in the form of an integral copy.

| Finding No 1 |  |  |  |  |
| --- | --- | --- | --- | --- |
| Description of nonconformity |  |  |  |  |
| The internal audit report lacks adequate process coverage and objective evidence : - Only one person (Quality Director) was audited, limiting verification of QMS implementation across functions. Furthermore, key processes were not audited, including supplier control, complaint handling, software validation, and internal audits. - The report does not provide adequate objective evidence (e.g., sampled records, checklists) to support conformity claims. - The report and the agenda do not have identification as part of the QMS system |  |  |  |  |
| Requirement |  |  |  |  |
| ISO 13485:2016 8.2.4 MDR Annex IX 2.2 |  |  |  |  |
| Objective evidence of nonconformity |  |  |  |  |
| Internal Audit report |  |  |  |  |
| Classification |  |  |  |  |
| ☐ major NC | ☒ minor NC |  | ☐ observation |  |
| Auditor VC |  | Date 15/7/2025 |  |  |
| Client response |  |  |  | Date: |
| Root cause analysis |  |  |  |  |
| The internal audit was carried out by only one person, which limited the coverage and objectivity of the process. Furthermore, due to partial planning, some relevant processes were not included. The absence of objective evidence and the lack of formal identification in the documentation are due to insufficient application of the current audit procedure and the use of outdated templates. |  |  |  |  |
| Correction(s) to eliminate the nonconformity |  |  |  |  |
| Record the non-conformity in the quality system and provide traceability of the follow-up. |  |  |  |  |
| Corrective action(s) to eliminate the cause of the nonconformity |  |  |  |  |
| Review and update the internal audit procedure to: • Include minimum criteria for process coverage and evidence sampling, requiring all departments affected by the QMS to be audited. • Highlight the need to use controlled templates with identification in accordance with the document management system. • Ensure that future internal audits adequately plan the scope and provide for the review of key processes. |  |  |  |  |
| Submitted Evidence |  |  |  |  |
| Estimated timeframe for completion (if not already completed) |  |  |  |  |
| 31/10/2025 |  |  |  |  |
| HTCert review |  |  |  |  |
| ☐ acceptable |  |  |  |  |
| ☐ not acceptable, see description of remaining deficiency |  |  |  |  |
| Auditor |  | Date |  |  |

| Finding No 2 |  |  |  |  |
| --- | --- | --- | --- | --- |
| Description of nonconformity |  |  |  |  |
| The change handling SOP does not address impacts on MDR compliance, risk management, UDI-DI, or reporting to the Notified Body. |  |  |  |  |
| Requirement |  |  |  |  |
| MDR Annex IX 2.2, 2.4 |  |  |  |  |
| Objective evidence of nonconformity |  |  |  |  |
| PNT-005 |  |  |  |  |
| Classification |  |  |  |  |
| ☐ major NC | ☒ minor NC |  | ☐ observation |  |
| Auditor PM |  | Date 15/7/2025 |  |  |
| Client response |  |  |  | Date: |
| Root cause analysis |  |  |  |  |
| The change management procedure does not explicitly address the assessment of the impact on compliance with the MDR, nor on critical elements such as risk management, UDI-DI or the need for notification to the Notified Body. This omission is due to the fact that the procedure was initially adapted from a generic model without integrating all the specific requirements applicable to medical devices under the MDR, especially at an early stage of quality system development. |  |  |  |  |
| Correction(s) to eliminate the nonconformity |  |  |  |  |
| Record the non-conformity in the quality system and provide traceability of the follow-up. |  |  |  |  |
| Corrective action(s) to eliminate the cause of the nonconformity |  |  |  |  |
| Review and update the change control procedure to: • Include a specific section on regulatory impact assessment (MDR). • Add fields in the change request template to record the risk management impact analysis, UDI identifiers, technical documentation and the need for notification to HTCert. • Establish a checkpoint to review whether the change affects information already assessed by the Notified Body. |  |  |  |  |
| Submitted Evidence |  |  |  |  |
| Estimated timeframe for completion (if not already completed) |  |  |  |  |
| 31/10/2025 |  |  |  |  |
| HTCert review |  |  |  |  |
| ☐ acceptable |  |  |  |  |
| ☐ not acceptable, see description of remaining deficiency |  |  |  |  |
| Auditor |  | Date |  |  |

| Finding No 3 |  |  |  |  |
| --- | --- | --- | --- | --- |
| Description of nonconformity |  |  |  |  |
| The methodology used in the FMEA report for calculating the Risk Priority Number (RPN) is based on the multiplication of severity, probability, and detectability and does not align with the Risk Management procedure PNT-002.04, which defines the RPN calculation as the product of severity and probability only |  |  |  |  |
| Requirement |  |  |  |  |
| ISO13485 4.1.2b MDR Annex IX 2.2 |  |  |  |  |
| Objective evidence of nonconformity |  |  |  |  |
| MR-001.01_24 PNT-002 |  |  |  |  |
| Classification |  |  |  |  |
| ☐ major NC | ☒ minor NC |  | ☐ observation |  |
| Auditor VC |  | Date 15/7/2025 |  |  |
| Client response |  |  |  | Date: |
| Root cause analysis |  |  |  |  |
| A discrepancy has been identified between the methodology applied in the FMEA report and the risk management procedure PNT-002.04. While the procedure defines the calculation of the RPN as the multiplication of severity and probability (RPN = S × P), the FMEA report uses the traditional formula S × P × D (detectability). This inconsistency is due to the lack of alignment between the template used for the FMEA and the internal criteria defined in the procedure. |  |  |  |  |
| Correction(s) to eliminate the nonconformity |  |  |  |  |
| Record the non-conformity in the quality system and provide traceability of the follow-up. |  |  |  |  |
| Corrective action(s) to eliminate the cause of the nonconformity |  |  |  |  |
| • Updating the standard FMEA analysis template to correctly reflect the internal methodology for calculating the RPN (S × P). • Review of previous risk analyses to identify possible inconsistent documents. |  |  |  |  |
| Submitted Evidence |  |  |  |  |
| Estimated timeframe for completion (if not already completed) |  |  |  |  |
| 31/10/2025 |  |  |  |  |
| HTCert review |  |  |  |  |
| ☐ acceptable |  |  |  |  |
| ☐ not acceptable, see description of remaining deficiency |  |  |  |  |
| Auditor |  | Date |  |  |

| Finding No 4 |  |  |  |  |
| --- | --- | --- | --- | --- |
| Description of nonconformity |  |  |  |  |
| The issues related with the UDI (basic udi-di, udi-pi etc) are not clearly addressed in the QMS. (E.g. It is not covered the impact of versioning to the udi-pi, or the impact of change of the intended purpose of the product to the basic udi-di) |  |  |  |  |
| Requirement |  |  |  |  |
| MDR Article 27 |  |  |  |  |
| Objective evidence of nonconformity |  |  |  |  |
| PNT-019 LABELING MANAGEMENT |  |  |  |  |
| Classification |  |  |  |  |
| ☐ major NC | ☒ minor NC |  | ☐ observation |  |
| Auditor PM |  | Date 15/7/2025 |  |  |
| Client response |  |  |  | Date: |
| Root cause analysis |  |  |  |  |
| The lack of clear development on UDI-related aspects (including Basic UDI-DI, UDI-DI and UDI-PI) within the quality management system is due to a limited initial interpretation of the MDR requirements in this area. In particular, the impacts of certain changes on the unique identifier, such as changes in the version/revision of the software (UDI-PI) or changes in the intended use of the product (Basic UDI-DI), are not adequately addressed. |  |  |  |  |
| Correction(s) to eliminate the nonconformity |  |  |  |  |
| Record the non-conformity in the quality system and provide traceability of the follow-up. |  |  |  |  |
| Corrective action(s) to eliminate the cause of the nonconformity |  |  |  |  |
| Modification of the labelling control procedure to include: • Clear definition of UDI-DI, UDI-PI and Basic UDI-DI elements • Records and traceability of UDI code assignment. • Specification of the person responsible for generating, maintaining and communicating UDI codes. Update of the change control procedure to: • Include UDI impact analysis for every modification. • Document when a release generates a new UDI-PI. • Include criteria for determining when a change affects the Basic UDI-DI (e.g. change in intended use or product family). |  |  |  |  |
| Submitted Evidence |  |  |  |  |
| Estimated timeframe for completion (if not already completed) |  |  |  |  |
| 31/10/2025 |  |  |  |  |
| HTCert review |  |  |  |  |
| ☐ acceptable |  |  |  |  |
| ☐ not acceptable, see description of remaining deficiency |  |  |  |  |
| Auditor |  | Date |  |  |

| Finding No 5 |  |  |  |  |
| --- | --- | --- | --- | --- |
| Description of nonconformity |  |  |  |  |
| The SOP PNT-030 for Strategy for regulatory compliance does not clearly describes the identification of relevant legal requirements, the identification of the sources as well as the responsibilities for the tasks. In addition the follow up of CS and/or harmonized standards is not covered |  |  |  |  |
| Requirement |  |  |  |  |
| MDR Annex IX 2.2 |  |  |  |  |
| Objective evidence of nonconformity |  |  |  |  |
| PNT-030 REGULATORY MONITORING |  |  |  |  |
| Classification |  |  |  |  |
| ☐ major NC | ☒ minor NC |  | ☐ observation |  |
| Auditor PM |  | Date 15/7/2025 |  |  |
| Client response |  |  |  | Date: |
| Root cause analysis |  |  |  |  |
| Procedure PNT-030 does not explicitly and in a structured way address the mechanisms to: • Identify applicable legal requirements (MDR, national legislation, etc.). • Determine the official sources of information (EU Official Journal, European Commission website, notified bodies, industry associations). • Assign clear responsibilities for monitoring these requirements. • Monitor changes in Common Specifications (CS) or relevant harmonised standards. • This omission is due to the fact that the procedure was initially drafted early in the development of the system, and was not revised after the full entry into force of the MDR. |  |  |  |  |
| Correction(s) to eliminate the nonconformity |  |  |  |  |
| Record the non-conformity in the quality system and provide traceability of the follow-up. |  |  |  |  |
| Corrective action(s) to eliminate the cause of the nonconformity |  |  |  |  |
| Revision of procedure PNT-030, including: • Specific section for the identification of applicable legal and regulatory requirements, including MDR, applicable directives, national legislation, and MDCG guidelines. • Inclusion of a list of official sources of updated information (websites, bulletins, standards databases, etc.). • Clear assignment of responsible parties for monitoring, reviewing and updating these requirements. • Description of the monitoring process for harmonised standards and Common Specifications (CS), including frequency, monitoring tools (e.g. CEN notifications, OJEU) and criteria for assessing their applicability. • Integration with the risk assessment and change management procedure to ensure that the system is updated when requirements change. |  |  |  |  |
| Submitted Evidence |  |  |  |  |
| Estimated timeframe for completion (if not already completed) |  |  |  |  |
| 31/10/2025 |  |  |  |  |
| HTCert review |  |  |  |  |
| ☐ acceptable |  |  |  |  |
| ☐ not acceptable, see description of remaining deficiency |  |  |  |  |
| Auditor |  | Date |  |  |

| Finding No 6 |  |  |  |  |
| --- | --- | --- | --- | --- |
| Description of nonconformity |  |  |  |  |
| No mechanism identified in the QMS, related to the identification of the official languages of the member states |  |  |  |  |
| Requirement |  |  |  |  |
| Objective evidence of nonconformity |  |  |  |  |
| PNT-001 DOCUMENT AND ARCHIVE MANAGEMENT |  |  |  |  |
| Classification |  |  |  |  |
| ☐ major NC | ☐ minor NC |  | ☒ observation |  |
| Auditor VC |  | Date 15/7/2025 |  |  |
| Client response |  |  |  | Date: |
| Root cause analysis |  |  |  |  |
| Correction(s) to eliminate the nonconformity |  |  |  |  |
| Corrective action(s) to eliminate the cause of the nonconformity |  |  |  |  |
| Submitted Evidence |  |  |  |  |
| Estimated timeframe for completion (if not already completed) |  |  |  |  |
| HTCert review |  |  |  |  |
| ☐ acceptable |  |  |  |  |
| ☐ not acceptable, see description of remaining deficiency |  |  |  |  |
| Auditor |  | Date |  |  |

| Finding No 7 |  |  |  |  |
| --- | --- | --- | --- | --- |
| Description of nonconformity |  |  |  |  |
| The criteria used for the supplier evaluation are not clearly stated. |  |  |  |  |
| Requirement |  |  |  |  |
| ISO13485 7.4.1 |  |  |  |  |
| Objective evidence of nonconformity |  |  |  |  |
| PNT-012 SELECTION, EVALUATION AND APPROVAL OF SUPPLIERS |  |  |  |  |
| Classification |  |  |  |  |
| ☐ major NC | ☒ minor NC |  | ☐ observation |  |
| Auditor VC |  | Date 15/7/2025 |  |  |
| Client response |  |  |  | Date: |
| Root cause analysis |  |  |  |  |
| The documentation of the supplier evaluation procedure lacks objective and specific criteria, which creates ambiguity in the evaluation and re-evaluation of suppliers. This is due to deficiencies in the definition of the procedure within the quality management system. |  |  |  |  |
| Correction(s) to eliminate the nonconformity |  |  |  |  |
| Record the non-conformity in the quality system and provide traceability of the follow-up. |  |  |  |  |
| Corrective action(s) to eliminate the cause of the nonconformity |  |  |  |  |
| Review and update the supplier evaluation procedure to include clearly defined, measurable criteria aligned with MDR requirements for the evaluation and re-evaluation of suppliers. Implement a standardized template to document each supplier evaluation and re-evaluation using the aforementioned criteria. |  |  |  |  |
| Submitted Evidence |  |  |  |  |
| Estimated timeframe for completion (if not already completed) |  |  |  |  |
| 31/10/2025 |  |  |  |  |
| HTCert review |  |  |  |  |
| ☐ acceptable |  |  |  |  |
| ☐ not acceptable, see description of remaining deficiency |  |  |  |  |
| Auditor |  | Date |  |  |