Annex 2_VERSION_05

- CAPA Report

| LAYER REPORT | CODE |
| --- | --- |
|  | AC-002/25 |

| Person opening the action: | Iván Pérez |  |  |  |  |  |  | Department |  |  |  | Quality Assurance |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Detection date | 15/07/2025 |  |  |  |  |  |  | Origin |  | Notified Body Audit |  |  |
| Description of the finding | The change handling SOP does not address impacts on MDR compliance, risk management, UDI-DI, or reporting to the Notified Body. |  |  |  |  |  |  |  |  |  |  |  |
| Root cause analysis | The change management procedure does not explicitly address the assessment of the impact on compliance with the MDR, nor on critical elements such as risk management, UDI-DI or the need for notification to the Notified Body. This omission is due to the fact that the procedure was initially adapted from a generic model without integrating all the specific requirements applicable to medical devices under the MDR, especially at an early stage of quality system development. |  |  |  |  |  |  |  |  |  |  |  |
| Immediate correction | Record the non-conformity in the quality system and provide traceability of the follow-up. |  |  |  |  |  |  |  |  |  |  |  |
| Corrective/preventive action | Review and update the change control procedure to: • Include a specific section on regulatory impact assessment (MDR). • Add fields in the change request template to record the risk management impact analysis, UDI identifiers, technical documentation and the need for notification to HTCert. • Establish a checkpoint to review whether the change affects information already assessed by the Notified Body. |  |  |  |  |  |  |  |  |  |  |  |
| Verification of the impact of the action on the product |  |  |  |  |  |  | Yes (if you check this box, explain below) |  |  |  | No |  |
|  | Is there an impact on regulatory requirements? |  |  |  |  |  | ☐ |  |  |  | ☒ |  |
|  | Is there an impact on safety and operation? |  |  |  |  |  | ☐ |  |  |  | ☒ |  |
|  | Impact assessment: |  |  |  |  |  | N/A |  |  |  |  |  |
| Responsible for the action | Iván Pérez |  |  |  |  |  |  |  | Department |  |  | Quality Assurance |
| Deadline | 31/10/2025 |  |  |  |  |  |  |  |  |  |  |  |
| To be completed by the person who opens | Signature and date |  |  |  |  |  |  |  |  |  |  |  |
| To be completed by the quality manager: | Signature and date Date: |  |  |  |  |  |  |  |  |  |  |  |
| Evaluation of the effectiveness of the change – To be completed when the change is made |  |  |  |  |  |  |  |  |  |  |  |  |
| Has the change been implemented successfully? |  | ☐ |  | Yes |  |  |  |  |  |  |  |  |
|  |  | ☐ |  | No | Observations |  |  |  |  |  |  |  |
| Effectiveness of change – To be completed when the change is made |  |  |  |  |  |  |  |  |  |  |  |  |
| Closing carried out by: |  |  |  |  |  | Date |  |  |  |  |  |  |