Annex 2_VERSION_05

- CAPA Report

| LAYER REPORT | CODE |
| --- | --- |
|  | AC-004/25 |

| Person opening the action: | Iván Pérez |  |  |  |  |  |  | Department |  |  |  | Quality Assurance |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Detection date | 15/07/2025 |  |  |  |  |  |  | Origin |  | Notified Body Audit |  |  |
| Description of the finding | The issues related with the UDI (basic udi-di, udi-pi etc) are not clearly addressed in the QMS. (E.g. It is not described the impact of versioning to the udi-pi, or the impact of change of the intended purpose of the product to the basic udi-di) |  |  |  |  |  |  |  |  |  |  |  |
| Root cause analysis | The lack of clear development on UDI-related aspects (including Basic UDI-DI, UDI-DI and UDI-PI) within the quality management system is due to a limited initial interpretation of the MDR requirements in this area. In particular, the impacts of certain changes on the unique identifier, such as changes in the version/revision of the software (UDI-PI) or changes in the intended use of the product (Basic UDI-DI), are not adequately addressed. |  |  |  |  |  |  |  |  |  |  |  |
| Immediate correction | Record the non-conformity in the quality system and provide traceability of the follow-up. |  |  |  |  |  |  |  |  |  |  |  |
| Corrective/preventive action | Modification of the labelling control procedure to include: - Clear definition of UDI-DI, UDI-PI and Basic UDI-DI elements - Records and traceability of UDI code assignment. - Specification of the person responsible for generating, maintaining and communicating UDI codes. Update of the change control procedure to: - Include UDI impact analysis for every modification. - Document when a release generates a new UDI-PI. - Include criteria for determining when a change affects the Basic UDI-DI (e.g. change in intended use or product family). |  |  |  |  |  |  |  |  |  |  |  |
| Verification of the impact of the action on the product |  |  |  |  |  |  | Yes (if you check this box, explain below) |  |  |  | No |  |
|  | Is there an impact on regulatory requirements? |  |  |  |  |  | ☒ |  |  |  | ☐ |  |
|  | Is there an impact on safety and operation? |  |  |  |  |  | ☐ |  |  |  | ☒ |  |
|  | Impact assessment: |  |  |  |  |  | The identified shortcoming has not had a direct impact on the product or its safety. However, it does represent a partial non-compliance with the MDR regarding the management of the UDI. The corrective actions put in place ensure that the UDI and its updates are now adequately addressed and documented in accordance with regulatory requirements. |  |  |  |  |  |
| Responsible for the action | Iván Pérez |  |  |  |  |  |  |  | Department |  |  | Quality Assurance |
| Deadline | 31/10/2025 |  |  |  |  |  |  |  |  |  |  |  |
| To be completed by the person who opens | Signature and date |  |  |  |  |  |  |  |  |  |  |  |
| To be completed by the quality manager: | Signature and date Date: |  |  |  |  |  |  |  |  |  |  |  |
| Evaluation of the effectiveness of the change – To be completed when the change is made |  |  |  |  |  |  |  |  |  |  |  |  |
| Has the change been implemented successfully? |  | ☐ |  | Yes |  |  |  |  |  |  |  |  |
|  |  | ☐ |  | No | Observations |  |  |  |  |  |  |  |
| Effectiveness of change – To be completed when the change is made |  |  |  |  |  |  |  |  |  |  |  |  |
| Closing carried out by: |  |  |  |  |  | Date |  |  |  |  |  |  |