## Frontpage
|  | VR-CARDIO Risk Analysis | Doc Code: MR-003.01 Version: 02 Date: 04/06/2025
Risk Management Team
Fernando Pozo (Executive Director) Iván Pérez (Quality Guarantee) José Martín (IT & Robotics) Vicente Gallego (Investigation and development) Raul Cordón (Sales) José Ángel Cabrera (Health Professional) Celeste Oliveira (PRRC) See Risks Management team separator for more details.
Summary
Based on the established risk management procedure, having carried out an in-depth analysis of the risks associated with VR-CARDIO Family and implementing appropriate control actions, we can conclude that the risks are now under control. The risk management measures implemented have effectively reduced the potential risks associated with the use of the VR-CARDIO-01 and VR-CARDIO-02 References.
Approved by: | Fernando Pozo
Date: 04/06/2025

## Hazard category
Hazard category | Hazard sub-category | Applicability | Possible contributing factors | Possible harms | Associated severity (see severity score)
Energy | Acoustic energy | Not applicable | infrasound sound pressure ultrasonic | __ | __
| Electrical energy | Not applicable | Electric fields Leakage current Magnetic fields Static discharge Voltage | __ | __
|  |  |  | __ | __
| Potential energy | Not applicable | — bending — compression — cutting, shearing — gravitational pull — suspended mass — tension — torsion | __ | __
| Mechanical energy | Not applicable | Kinetic energy — high pressure fluid injection — moving parts | __ | __
| Radiation energy | Not applicable | Ionizing radiation Non-ionizing radiation | __ | __
| Thermal energy |  | Cryogenic effects Hyperthermic effects | __ | __
Biological and chemical | Biological | Not applicable | Bacteria Fungi Parasites Prions Toxins Viruses | __ | __
| Chemical | Not applicable | Additives or processing aids Sterilization substantes Degradation products Particles (including micro- and nanoparticles) Pyrogenic Solvents Toxic | __ | __
| Biocompatibility | Not applicable | Allergenic — antiseptic substances — latex Immunosuppressive Irritants — cleaning residues Sensitizing | __ | __
Performance | Functionality | Applicable | Incorrect or inappropriate output or functionality Loss or deterioration of function | Impossibility to use the device | Minor. Need to repeat the procedure.
| Use error | Applicable | Attentional failure (human error risk) | Incorrect use of the device | Minor. Need to repeat the procedure.
| Operating instructions | Applicable | Inadequate specification of accessories to be used with the medical device Over complicated instructions for use | Incorrect use of the device | Serious. Misdiagnosis leading to inappropriate treatment or treatment delays
|  |  |  | Impossibility to use the device | Minor. Need to repeat the procedure.
| Warnings | Applicable | Warnings of side effects Warnings of hazards related to the device | Incorrect use of the device | Serious. Misdiagnosis leading to inappropriate treatment or treatment delays
| Data | Applicable | — access — availability — confidentiality — data transfer — integrity | Incorrect use of the device | Serious. Misdiagnosis leading to inappropriate treatment or treatment delays
|  |  |  | Incorrect use of the device - (data privacy and security breaches) | Minor. Unauthorized access to patient data
|  |  |  | Impossibility to use the device | Minor. Need to repeat the procedure.
| Diagnostic information | Applicable | — examination result — image artefacts — image orientation — image resolution — patient identity / information | Incorrect use of the device (for example, misinterpretation of the outputs of the device) | Serious. Misdiagnosis leading to inappropriate treatment or treatment delays
|  |  |  | Incorrect output of the device (for example, bad quality images, poor resolution) | Serious. Misdiagnosis leading to inappropriate treatment or treatment delays
|  |  |  | Impossibility to use the device | Minor. Need to repeat the procedure.

## Risk Matrix
Definition of risk | [5 empty] | Risk assessment |  |  | Risk acceptance | Risk control (if applicable) | Risk Control (if applicable) |  |  | Risk acceptance | Additional measures | Residual Risk Evaluation
# | Process/system failure | Phase | Potential Hazard | Hazardous situation | Potential harm | Probability | Severity | RPN = P x S | Acceptable? | Risk Control Measures | Probability | Severity | RPN = P x S | Acceptable? | Create new risks? | Acceptable?
VR-CARDIO Family
1 | Activation/Installation/Configuration Process | Software Implementation | Authentication and authorization issues | Users are incorrectly granted or denied access due to configuration or logic flaws. | System instability, security vulnerabilities | 4 | 3 | 12 | Tolerable | Define access control roles and implement secure configuration templates. Conduct periodic audits of permission assignments. Establish approval workflows for privilege changes. | 1 | 3 | 3 | Acceptable | No | Yes
2 | Design and development | Software Requirements Specification | Ambiguous requirements | Requirements are vague or open to interpretation during design and coding. | Misdiagnosis leading to inappropriate treatment logic | 3 | 3 | 9 | Tolerable | Conduct joint review sessions with users and developers. Define a controlled vocabulary for clinical and system terms. Maintain traceable requirement-to-design linkage. | 1 | 3 | 3 | Acceptable | No | Yes
3 | Design and development | Software Requirements Specification | Incomplete requirements | Developers lack complete information to cover all use cases. | Misdiagnosis due to gaps in functionality | 3 | 3 | 9 | Tolerable | Host structured workshops to validate requirements with end users. Document stakeholder feedback and perform gap analysis. | 1 | 3 | 3 | Acceptable | No | Yes
4 | Design and development | Software Requirements Specification | Unrealistic expectations | The project team commits to infeasible features or timelines. | Delayed product release, reduced software quality | 3 | 2 | 6 | Tolerable | Establish realistic delivery milestones based on resource availability. Maintain a documented project expectation register. | 1 | 2 | 2 | Acceptable | No | Yes
5 | Design and development | Software Design and Architecture | Inadequate data modeling | Data structures fail to reflect user workflow or clinical logic. | Development delays, inability to scale | 2 | 2 | 4 | Acceptable | Define standards for schema documentation. Validate model assumptions with clinical experts. | 1 | 2 | 2 | Acceptable | No | Yes
6 | Design and development | Software Design and Architecture | Inadequate documentation | System behavior or structure is unclear to developers or users. | Maintenance gaps, errors during updates | 3 | 3 | 9 | Tolerable | Establish documentation templates for each lifecycle stage. Perform quarterly documentation reviews. Train teams on documentation update protocols. | 1 | 3 | 3 | Acceptable | No | Yes
7 | Design and development | Software Design and Architecture | Incomplete documentation | Missing instructions cause confusion in design or testing. | Design misalignment with clinical needs | 3 | 3 | 9 | Tolerable | Gather feedback from testers and clinicians. Track documentation completeness using controlled checklists. Include review checkpoints in design validation. | 1 | 3 | 3 | Acceptable | No | Yes
8 | Design and development | Software Design and Architecture | Lack of version control for designs | Design artifacts diverge without tracking. | Design deviations, scope changes | 3 | 2 | 6 | Tolerable | Apply version control to all design artifacts using a centralized repository. Require version tags before implementation. Include design approval workflows as part of change management. | 1 | 2 | 2 | Acceptable | No | Yes
9 | Design and development | Software Design and Architecture | Poorly defined system requirements | Requirements are not specific enough for safe implementation. | Performance issues, software instability | 3 | 3 | 9 | Tolerable | Clarify and validate requirements through user involvement. Implement formal review cycles. | 1 | 3 | 3 | Acceptable | No | Yes
10 | Design and development | Software Design and Architecture | Lack of accessibility considerations | System is not usable by users with assistive tech or visual limitations. | High maintenance overhead, loss of inclusivity | 3 | 3 | 9 | Tolerable | Emphasize modular design principles, conduct code reviews, and encourage the development of independent, reusable modules. | 1 | 3 | 3 | Acceptable | No | Yes
11 | Design and development | Software Design and Architecture | Lack of modularity | Changes in one area affect others unexpectedly. | Integration issues, system instability | 3 | 3 | 9 | Tolerable | Apply modular architecture principles, isolate critical functions into reusable components, and enforce interface contracts between modules. | 1 | 3 | 3 | Acceptable | No | Yes
12 | Design and development | Software Design and Architecture | Overly complex design | User flow becomes difficult to understand or use. | User confusion, high support burden | 1 | 3 | 3 | Acceptable | Establish and enforce design guidelines, conduct periodic reviews, and provide training on design patterns. | 1 | 3 | 3 | Acceptable | No | Yes
13 | Design and development | Software Design and Architecture | Lack of consistency in design | Repetitive or inconsistent interface logic. | User frustration, increased support requests | 1 | 3 | 3 | Acceptable | Promote interface clarity through design reviews, user feedback, and usability testing. | 1 | 3 | 3 | Acceptable | No | Yes
14 | Design and development | Software development | Unclear interface design | Ambiguous or non-standard UI elements. | Implementation errors, training complexity | 3 | 3 | 9 | Tolerable | Apply user-centered design practices, perform heuristic evaluations, and clarify controls through guided workflows and tooltips. | 1 | 3 | 3 | Acceptable | No | Yes
15 | Design and development | Software development | Code dependencies | Modules break due to uncontrolled dependency upgrades. | System crashes, data loss | 3 | 3 | 9 | Tolerable | Lock dependency versions, implement regression test suites, and review external dependency impact at each release cycle. | 1 | 3 | 3 | Acceptable | No | Yes
16 | Design and development | Software development | Lack of code reviews | Unchecked code with logic errors enters production. | Decreased code quality, regression risk | 3 | 2 | 6 | Tolerable | Establish peer-review policies, use pull request templates, and track review metrics across teams. | 2 | 2 | 4 | Acceptable | No | Yes
17 | Design and development | Software development | Inadequate code documentation | Developers cannot understand legacy or third-party code. | Maintenance risk, development delays | 2 | 2 | 4 | Acceptable | Emphasize thorough code documentation, provide guidelines, and perform regular checks to ensure code is well documented. | 1 | 2 | 2 | Acceptable | No | Yes
18 | Design and development | Software development | Lack of collaboration between development teams | Parallel development leads to incompatible code branches. | Merge conflicts, software bugs | 2 | 2 | 4 | Acceptable | Use shared documentation tools, hold weekly alignment meetings, and track integration milestones cross-functionally. | 1 | 2 | 2 | Acceptable | No | Yes
19 | Design and development | Software development | Insufficient training in new technologies | Team lacks familiarity with new frameworks or platforms. | Coding inefficiencies, bug introduction | 2 | 2 | 4 | Acceptable | Conduct onboarding programs for new stacks, provide sandbox environments, and encourage technical mentorship. | 1 | 2 | 2 | Acceptable | No | Yes
20 | Production | Software Production and Release | Dependency management issues | Packages are outdated or cause version conflicts. | Compatibility issues, feature gaps | 3 | 3 | 9 | Tolerable | Implement model drift monitoring, retrain models regularly, and address changes in data distribution to maintain model accuracy. | 1 | 3 | 3 | Acceptable | No | Yes
21 | Production | Software Production and Release | Code integration failures | Code changes fail during merge or build process. | Implementation failures, system instability | 3 | 3 | 9 | Tolerable | Establish integration branch testing, apply continuous integration practices | 1 | 3 | 3 | Acceptable | No | Yes
22 | Production | Software Production and Release | Downtime due to server failure | Backend server is unavailable. | Data loss, diagnostic service interruption | 3 | 3 | 9 | Tolerable | Perform tests, and implement automated integration checks. | 1 | 3 | 3 | Acceptable | No | Yes
23 | Production | Software update | Use of unauthorized software | External or unvalidated tools compromise the system. | Equipment malfunction, unpredictable behavior | 2 | 3 | 6 | Tolerable | Maintain an approved software inventory, implement whitelist policies, and validate 3rd-party tools before use. | 1 | 3 | 3 | Acceptable | No | Yes
24 | Technical Processes | Software Production and Release | Insufficient load testing | System fails under real-world traffic. | Performance bottlenecks, diagnostic delay | 2 | 2 | 4 | Acceptable | Simulate concurrent usage with realistic load patterns, and benchmark thresholds using tools like JMeter or k6. | 1 | 2 | 2 | Acceptable | No | Yes
25 | Technical Processes | Technical Support and Maintenance | Improper software patching | Partial patches break consistency across system components. | Security issues, instability | 2 | 2 | 4 | Acceptable | Automate patch deployment pipelines, verify patch integrity, and maintain patch logs with change descriptions. | 1 | 2 | 2 | Acceptable | No | Yes
26 | Technical Processes | Technical Support and Maintenance | Database performance degradation | Queries run slowly or time out. | Reduced system responsiveness | 2 | 2 | 4 | Acceptable | Optimize indexing strategies, implement automated query tuning, and conduct load/performance tests during release cycles. | 1 | 2 | 2 | Acceptable | No | Yes
27 | Technical Processes | Technical Support and Maintenance | Inadequate documentation maintenance | Documentation is outdated or inaccurate. | Errors during updates, support inefficiency | 2 | 3 | 6 | Tolerable | Integrate doc updates into release process, assign documentation ownership, and schedule quarterly documentation audits. | 1 | 3 | 3 | Acceptable | No | Yes
28 | Technical Processes | Technical Support and Maintenance | Insufficient backup and recovery measures | No viable recovery path after a crash or corruption. | Permanent data loss | 2 | 2 | 4 | Acceptable | Establish automated, encrypted backup system with offsite storage. Perform regular backup validation drills. | 1 | 2 | 2 | Acceptable | No | Yes
29 | Cybersecurity | Violations of data retention policy | Security breach | Data stolen due to software weakness. | Legal exposure | 2 | 2 | 4 | Acceptable | Conduct penetration testing, implement encryption at rest/in transit, and train team on secure coding standards. | 2 | 2 | 4 | Acceptable | No | Yes
30 | Cybersecurity | Inadequate data anonymization | Inadequate encryption | Encrypted channels are misconfigured or weak. | Privacy breaches | 4 | 2 | 8 | Tolerable | Use modern encryption protocols. Implement regular key rotation and vulnerability scanning. | 2 | 2 | 4 | Acceptable | No | Yes
31 | Cybersecurity | Lack of clear communication | Inadequate user communication practices | Users receive unclear or inconsistent system instructions. | User confusion, incorrect usage, diagnostic errors | 2 | 3 | 6 | Tolerable | Revise user documentation for clarity. Include visuals and action-based guides. Test content with intended users. | 1 | 3 | 3 | Acceptable | No | Yes
32 | Cybersecurity | Lack of user consent management | Inadequate consent practices | System collects or uses patient data without formalized user consent. | Legal risk, privacy violations, regulatory sanctions | 2 | 2 | 4 | Acceptable | Require consent acknowledgment during onboarding. Maintain consent logs and include privacy policy in system workflow. | 1 | 2 | 2 | Acceptable | No | Yes
33 | Cybersecurity | Unauthorized access | Weak authentication or authorization | Unauthorized users can access the system or authorized users are blocked. | Data breaches, denial of service, compromised system integrity | 2 | 2 | 4 | Acceptable | Enforce strong password rules and two-factor authentication. Periodically review user permissions. | 1 | 2 | 2 | Acceptable | No | Yes
34 | Cybersecurity | Unauthorized data sharing | Insufficient controls for data sharing | Data is shared without proper access restriction or logging. | Privacy breaches, unauthorized disclosure, regulatory non-compliance | 4 | 2 | 8 | Tolerable | Control data export paths and require sharing permissions. Maintain export logs and enforce user-level restrictions. | 2 | 2 | 4 | Acceptable | No | Yes
35 | Cybersecurity | Lack of ongoing user privacy education | Infrequent education on user privacy practices | Users are unaware of best practices for handling sensitive health data. | Privacy violations due to mishandling of personal information | 2 | 2 | 4 | Acceptable | Schedule annual privacy training. Include onboarding modules and regular refresher quizzes. | 1 | 2 | 2 | Acceptable | No | Yes
36 | Cybersecurity | Violations of data protection regulations | Non-compliance with data protection regulations | Data is stored, processed, or transferred in violation of legal frameworks (e.g., GDPR). | Legal penalties, reputational damage, business continuity risks | 2 | 2 | 4 | Acceptable | Conduct Data Protection Impact Assessments (DPIAs). Implement internal audits and update policies regularly. | 1 | 2 | 2 | Acceptable | No | Yes
37 | Cybersecurity | Weak password policies | Inadequate password security policies | System allows weak, reused, or non-expiring passwords. | Account compromise, unauthorized system access | 4 | 2 | 8 | Tolerable | Require minimum complexity, block reuse, and auto-expire credentials. Provide secure password reset flow. | 2 | 2 | 4 | Acceptable | No | Yes
38 | User Management | User Management and Data Privacy | Data Retention Policy Violations | Data stored longer than legally allowed. | Legal consequences, privacy breaches | 2 | 2 | 4 | Acceptable | Implement retention rules in database, auto-purge expired data, and perform retention policy audits annually. | 1 | 2 | 2 | Acceptable | No | Yes
39 | User Management | User Management and Data Privacy | Insufficient data access controls | Unauthorized users access sensitive data. | Data breaches, reputation risk | 3 | 2 | 6 | Tolerable | Apply role-based access controls (RBAC), enable access logging, and review permissions quarterly. | 2 | 2 | 4 | Acceptable | No | Yes
40 | User Management | User Management and Data Privacy | Insufficient data anonymization | Identifiable patient info remains accessible. | Privacy breaches, GDPR violation | 3 | 2 | 6 | Tolerable | Apply structured anonymization before analysis, validate with test cases, and encrypt raw data at source. | 2 | 2 | 4 | Acceptable | No | Yes
41 | User Management | User Management and Data Privacy | Insufficient data backup | Failure occurs before backup is completed. | Data loss, interruption of care | 2 | 3 | 6 | Tolerable | Implement incremental backup scheduling, and set up alerts for incomplete backup cycles. | 1 | 3 | 3 | Acceptable | No | Yes
42 | User Management | User Management and Data Privacy | Insufficient data encryption | Data stored or transmitted in cleartext. | Data leakage, privacy breaches | 2 | 2 | 4 | Acceptable | Implement robust anonymization practices. | 2 | 2 | 4 | Acceptable | No | Yes
43 | User Management | User Management and Data Privacy | Lack of clear communication | Users do not understand software purpose or status. | User frustration, low adoption | 2 | 2 | 4 | Acceptable | Add real-time system status indicators, update tooltips, and clarify workflows through onboarding prompts. | 1 | 2 | 2 | Acceptable | No | Yes
44 | User Management | User Management and Data Privacy | Lack of user consent management | System captures or uses data without verified consent. | Legal risk, patient mistrust | 2 | 2 | 4 | Acceptable | Implement explicit consent prompts with opt-in controls, version consent forms, and log acceptance metadata. | 1 | 2 | 2 | Acceptable | No | Yes
45 | User Management | User Management and Data Privacy | Unauthorized access | Credentials reused or exposed. | Data breaches, privacy violations | 2 | 2 | 4 | Acceptable | Enforce unique user accounts, apply strong password policy, and enable multi-factor authentication (MFA). | 1 | 2 | 2 | Acceptable | No | Yes
46 | User Management | User Management and Data Privacy | Unauthorized data sharing | Data sent to unapproved third parties. | Legal exposure, patient harm | 3 | 2 | 6 | Tolerable | Set user-based data export permissions, encrypt shared files, and keep audit logs of shared sessions. | 2 | 2 | 4 | Acceptable | No | Yes
47 | User Management | User Management and Data Privacy | Lack of continuing education on user privacy | Staff mishandles data due to unawareness. | User error, increased risk of breaches | 2 | 2 | 4 | Acceptable | Provide mandatory training on privacy protocols, conduct scenario-based tests, and track training completion. | 1 | 2 | 2 | Acceptable | No | Yes
48 | User Management | User Privacy Compliance | Violations of data protection regulations | Data processed without lawful basis. | Fines, loss of regulatory trust | 2 | 2 | 4 | Acceptable | Implement data portability mechanisms, assess user requirements, and allow users to transfer their data between systems transparently. | 1 | 2 | 2 | Acceptable | No | Yes
49 | User Management | Security and User Access Control | Weak password policies | Short or easily guessed passwords allowed. | Unauthorized access, security breaches | 3 | 2 | 6 | Tolerable | Enforce complex passwords, expiration policies, and prohibit reuse of past credentials. | 2 | 2 | 4 | Acceptable | No | Yes
50 | Usability | User Training | Inappropriate instructions for use | IFU lacks clarity or completeness. | Users cannot operate system correctly | 2 | 2 | 4 | Acceptable | Conduct usability testing of IFU drafts, update IFU post-release. | 1 | 2 | 2 | Acceptable | No | Yes
51 | Usability | User Training | Inconsistent user onboarding | New users receive different information or training. | Misuse of tool, poor diagnostic outcomes | 2 | 2 | 4 | Acceptable | Develop and standardize an onboarding module, include tool walkthroughs, and track onboarding completion in the LMS. | 1 | 2 | 2 | Acceptable | No | Yes
52 | Usability | Login | Unable to authenticate | Login fails for technical or user reasons. | Blocked access to clinical system | 2 | 2 | 4 | Acceptable | Monitor authentication logs, enable support fallback option, and improve error messages to guide troubleshooting. | 1 | 2 | 2 | Acceptable | No | Yes
53 | Usability | Login | Authentication issues | Login fails for technical or user reasons. | Blocked access to clinical system | 2 | 2 | 4 | Acceptable | Add monitoring for auth endpoints, provide recovery mechanisms (e.g., password reset), and publish downtime alerts. | 1 | 2 | 2 | Acceptable | No | Yes
54 | Usability | Log out | User does not log out after using the software | Session remains open and unmonitored. | Unauthorized access by another user | 3 | 3 | 9 | Tolerable | Enforce auto logout after inactivity, show session expiry countdown, and log forced terminations. | 2 | 3 | 6 | Acceptable | No | Yes
55 | Usability | Task management | Non-intuitive interface design for new users | Users cannot locate critical features. | Inability to perform diagnostic tasks | 3 | 3 | 9 | Tolerable | Conduct usability tests with new users, adjust UI hierarchy based on feedback, and enhance tooltips for navigation. | 2 | 3 | 6 | Acceptable | No | Yes
56 | Usability | Access to recording tool | User inability to access their profile in vr-cardio explore | System won't load or recognize login session. | No diagnostic access | 2 | 2 | 4 | Acceptable | Improve session state handling, add retry logic for failed profile loads, and guide users through recovery path. | 1 | 2 | 2 | Acceptable | No | Yes
57 | Usability | Register a patient in the recording tool | User inability to register a patient | Data input or validation error prevents new patient entry. | No session can be initiated | 2 | 2 | 4 | Acceptable | Add input validation feedback, pre-fill standard fields where possible, and include fallback manual entry option. | 1 | 2 | 2 | Acceptable | No | Yes
58 | Usability | Data collection | User inability to record a data collection session | System fails to detect signal or save input. | Lost clinical data | 2 | 2 | 4 | Acceptable | Include signal quality indicators and automated recording status alerts. | 1 | 2 | 2 | Acceptable | No | Yes
59 | Usability | Data processing | Inability of the user to process the data collected | Data files are corrupted or unsupported. | No diagnostic output available | 3 | 2 | 6 | Tolerable | Define compatible data types. Include automated file validation before processing. | 1 | 2 | 2 | Acceptable | No | Yes
60 | User Handling | Device usage | Untrained user operation | Device used by individuals without prior training | Misuse of features, incorrect interpretation | 3 | 3 | 9 | Tolerable | Provide mandatory training prior to use. Restrict access until training is confirmed. Include verification in onboarding. | 1 | 3 | 3 | Acceptable | No | Yes
61 | User Handling | Device usage | IFU not consulted | User skips or misinterprets instructions | Improper use, inability to access features | 3 | 2 | 6 | Tolerable | Make reading the IFU part of the software setup. | 2 | 2 | 4 | Acceptable | No | Yes
62 | User Handling | Device usage | Use of incorrect external device size | Patient is connected to mismatched data acquisition system | Poor data capture, missed arrhythmia | 3 | 2 | 6 | Tolerable | List compatible device sizes in IFU. | 2 | 2 | 4 | Acceptable | No | Yes
63 | User Handling | Device usage | Credential sharing | Credentials are shared across users or clinics | Breach of sensitive patient data | 3 | 2 | 6 | Tolerable | Enforce single-login-per-session. Alert on multiple logins. | 2 | 2 | 4 | Acceptable | No | Yes
64 | User Handling | Device usage | Incorrect patient registration or signal acquisition | User enters wrong ID or poorly attaches electrodes | Inaccurate analysis, wrong patient file | 3 | 3 | 9 | Tolerable | Implement data verification prompts. Require patient ID confirmation before saving. | 1 | 3 | 3 | Acceptable | No | Yes
65 | User Handling | Device usage | Internet disconnection | Session disconnects mid-transmission | Data loss, session failure | 3 | 2 | 6 | Tolerable | Include autosave for session data. Warn users of connection loss. Add ability to resume interrupted sessions. | 1 | 2 | 2 | Acceptable | No | Yes
66 | User Handling | Device usage | Algorithm misinterpreted as final diagnosis | Users treat algorithm output as definitive | Missed or inappropriate clinical response | 3 | 3 | 9 | Tolerable | Display disclaimer prominently: "For diagnostic support only." Include “review required” flags. | 1 | 3 | 3 | Acceptable | No | Yes
67 | User Handling | Device usage | Heart maps interpreted in isolation | Clinicians rely solely on VR data | Misdiagnosis without clinical context | 3 | 3 | 9 | Tolerable | Add prompts requiring secondary confirmation before report export. Include contextual disclaimers. | 1 | 3 | 3 | Acceptable | No | Yes
68 | User Handling | Device usage | Filter misplacement | Incorrect filter applied to wrong data segment | Distorted results | 3 | 3 | 9 | Tolerable | Match filter ID with source data tags. Validate filters before display. | 1 | 3 | 3 | Acceptable | No | Yes
69 | User Handling | Device usage | Poor data acquisition | Signal quality too low to render panels accurately | Misleading or false visualization | 3 | 3 | 9 | Tolerable | Require reacquisition if signal falls below threshold. | 1 | 3 | 3 | Acceptable | No | Yes
VR-CARDIO Explore (Ref. VR-CARDIO-01) Specific Risks
70 | Activation/Installation/Configuration Process | Software Implementation | Incompatibility with certain browsers | User opens system in an unsupported browser, causing critical functionality to fail or behave unpredictably. | Loss of critical data, compromised system integrity | 4 | 3 | 12 | Tolerable | Perform browser compatibility testing across supported environments. Include clear browser usage instructions in the IFU. | 1 | 3 | 3 | Acceptable | No | Yes
71 | Activation/Installation/Configuration Process | Software Implementation | Data loss during deployment | Deployment occurs without backup or rollback, leading to partial or complete data loss. | Extended system downtime, user interruptions, loss of critical data, compromised system integrity | 3 | 3 | 9 | Tolerable | Develop and document rollback and backup procedures. Schedule regular simulations of rollback drills. | 1 | 3 | 3 | Acceptable | No | Yes
72 | Activation/Installation/Configuration Process | Software Implementation | Service interruptions during deployment | Users access the system during an incomplete update cycle. | Prolonged system instability, data corruption | 3 | 3 | 9 | Tolerable | Establish controlled deployment windows. Notify users prior to system maintenance. Implement automatic user session lockout during updates and monitor rollback logs for failures. | 1 | 3 | 3 | Acceptable | No | Yes
73 | Design and development | Software Design and Architecture | Insufficient cross-browser compatibility | Users on certain browsers encounter partial UI or broken components. | Downtime, data loss | 3 | 3 | 9 | Tolerable | Develop a comprehensive disaster recovery plan, including backup strategies, data recovery procedures, and system restore plans. | 1 | 3 | 3 | Acceptable | No | Yes
74 | Data processing | SOUP: Interacts with OS (os) | OS command errors | System executes unintended or malformed operating system commands. | Security vulnerabilities, unauthorized system access | 2 | 3 | 6 | Tolerable | Sanitize input, block direct shell commands where avoidable, and log command execution paths. | 1 | 3 | 3 | Acceptable | No | Yes
75 | Data processing | SOUP: Handles dates and times (datetime) | Date/time handling errors | Timestamps are incorrectly recorded or compared across modules. | Incorrect data correlation, flawed diagnostic timelines | 3 | 3 | 9 | Tolerable | Perform unit tests; validate date and time inputs and outputs | 1 | 3 | 3 | Acceptable | No | Yes
76 | Data processing | SOUP: Manages JSON data (json) | Incorrect JSON data handling | Data is parsed incorrectly, omitted, or misinterpreted. | Corrupted records, data loss, software crashes | 3 | 3 | 9 | Tolerable | Validate JSON schemas on both ends. Log and notify for malformed structures. | 1 | 3 | 3 | Acceptable | No | Yes
77 | Data processing | SOUP: Accesses interpreter (sys) | Interpreter argument errors | Improper arguments sent to script interpreters or compilers. | Program crash, unhandled exceptions, loss of functionality | 2 | 3 | 6 | Tolerable | Validate arguments and environment variables. Use defensive programming practices. | 1 | 3 | 3 | Acceptable | No | Yes
78 | Data processing | SOUP: Provides platform info (platform) | Platform information errors | System misidentifies or misreports platform parameters (e.g., OS version, GPU). | Compatibility failures, unstable execution | 2 | 3 | 6 | Tolerable | Validate platform information; Perform tests across different environments | 1 | 3 | 3 | Acceptable | No | Yes
79 | Data processing | SOUP: Executes system commands (subprocess) | Command execution errors | Execution of commands fails silently or causes unintended side effects. | Software instability, data corruption | 2 | 3 | 6 | Tolerable | Implement logging for all commands. Validate success/failure and implement rollback if needed. | 1 | 3 | 3 | Acceptable | No | Yes
80 | Data processing | SOUP: Provides time functions (time) | Time-related errors | Latency or timing mismatches disrupt processes requiring synchronization. | Signal distortion, diagnosis inconsistency | 2 | 3 | 6 | Tolerable | Validate time usage. Monitor latency and flag excessive delays. | 1 | 3 | 3 | Acceptable | No | Yes
81 | Data processing | SOUP: Math/stat calculations (numpy) | Calculation errors | Software produces mathematically incorrect outputs. | Inaccurate diagnosis, unreliable analytics | 3 | 3 | 9 | Tolerable | Unit test all formula paths. Validate against known test data and clinical tolerances. | 1 | 3 | 3 | Acceptable | No | Yes
82 | Data processing | SOUP: Data analysis (pandas) | Data manipulation errors | Programmatic operations alter data in unintended ways. | Diagnostic misinterpretation, data corruption | 3 | 3 | 9 | Tolerable | Apply strict input validation. Version control transformation scripts and run integrity checks. | 1 | 3 | 3 | Acceptable | No | Yes
83 | Data processing | SOUP: Sends HTTP requests (request) | HTTP request errors | Network communications between system components fail. | Data loss, broken features, unresponsive system | 3 | 3 | 9 | Tolerable | Retry on failure with backoff logic. Log failures and prompt users to retry or check network. | 1 | 3 | 3 | Acceptable | No | Yes
84 | Data processing | SOUP: Works with 3D meshes (trimesh) | 3D mesh handling errors | VR heart/tissue models fail to render or deform inaccurately. | Diagnostic support unavailable or misleading | 2 | 3 | 6 | Acceptable | Run integrity check on 3D files. Alert users of corrupt meshes and log failure types. | 1 | 3 | 3 | Acceptable | No | Yes
85 | Data processing | SOUP: Web framework (flask) | Web application errors | Web-based modules freeze, crash, or behave unpredictably. | Data breaches, patient care disruption | 3 | 3 | 9 | Tolerable | Conduct regression tests across supported browsers. Monitor front-end crashes and maintain SLA targets. | 1 | 3 | 3 | Acceptable | No | Yes
VR-CARDIO VHET360 (Ref. VR-CARDIO-02) Specific Risks
86 | Production | Software update | Software update failures | Update leaves system in non-functional state. | Equipment malfunction | 2 | 3 | 6 | Tolerable | Define rollback strategy, use staged deployments, and verify updates in test environments before release. | 1 | 3 | 3 | Acceptable | No | Yes
87 | Usability | Open visualization tool | User inability to open vrcardio-vhet360 | System fails to start VR environment. | Diagnostic feature unavailable | 2 | 2 | 4 | Acceptable | Implement startup diagnostic checks for VR modules. | 1 | 2 | 2 | Acceptable | No | Yes
88 | Usability | Enter in the visualization tool | User inability to log in to vrcardio-vhet360 | System fails to start VR environment. | Diagnostic feature unavailable | 2 | 2 | 4 | Acceptable | Add login diagnostics to detect service errors. Improve guidance on login prerequisites and fallback recovery steps. | 1 | 2 | 2 | Acceptable | No | Yes
89 | Usability | Observe case in the visualisation tool | User inability to upload data from a recorded session to vrcardio-vhet360 | System doesn't recognize source file or session. | Lost patient case, no analysis possible | 3 | 3 | 9 | Tolerable | Implement file validation pre-upload. | 2 | 3 | 6 | Acceptable | No | Yes
90 | Usability | Capturing patient data | The image does not accurately represent the injured heart area. | Visualization fails to map accurate lesion area. | Incorrect or delayed diagnosis | 3 | 3 | 9 | Tolerable | Validate imaging algorithms during verification. | 1 | 3 | 3 | Acceptable | No | Yes
91 | Usability | Capturing patient data | Tergiversation of results | Results are edited or presented in misleading ways. | Misdiagnosis | 2 | 3 | 6 | Tolerable | Implement audit trails for result generation and export. Protect diagnostic data integrity through tamper-evident logs. | 1 | 3 | 3 | Acceptable | No | Yes
92 | Usability | Capturing patient data | Insufficient visibility | Screen resolution or signal strength impairs view. | Diagnostic delay or misinterpretation | 3 | 3 | 9 | Tolerable | Provide hardware display minimum specs in IFU. Allow user control over brightness and contrast within the application. | 1 | 3 | 3 | Acceptable | No | Yes
93 | Usability | Capturing patient data | Incorrect potential mapping | Error in signal processing. | Wrong diagnosis of arrhythmia origin | 2 | 3 | 6 | Tolerable | Improve signal acquisition logic and validate mapping accuracy. | 1 | 3 | 3 | Acceptable | No | Yes
94 | Usability | Capturing patient data | Poor image resolution | Rendered images are too pixelated. | Diagnosis not possible | 2 | 3 | 6 | Tolerable | Apply resolution thresholds in rendering engine. Display alerts if device resolution is below minimum requirement. | 1 | 3 | 3 | Acceptable | No | Yes
95 | Usability | Capturing patient data | Incorrect image orientation | VR heart appears rotated or mirrored. | Misleading spatial interpretation | 2 | 3 | 6 | Tolerable | Implement an orientation check on launch and allow manual correction. | 1 | 3 | 3 | Acceptable | No | Yes
96 | Usability | Capturing patient data | Poor quality in obtained signals | Signal is distorted beyond interpretation. | Cannot detect arrhythmia | 3 | 3 | 9 | Tolerable | Apply real-time signal quality validation. Provide acquisition troubleshooting guidance. | 1 | 3 | 3 | Acceptable | No | Yes
97 | Usability | Capturing patient data | Excessive noise in obtained signals | Signal is distorted beyond interpretation. | Cannot detect arrhythmia | 3 | 3 | 9 | Tolerable | Introduce filtering presets and quality thresholds. Log signal quality at time of recording. | 1 | 3 | 3 | Acceptable | No | Yes
98 | Usability | Capturing patient data | Lack of biological signal in obtained signals | Heart signal capture fails completely. | No diagnostic basis | 3 | 3 | 9 | Tolerable | Display immediate alerts on acquisition failure. Guide user through connection or setup troubleshooting. | 1 | 3 | 3 | Acceptable | No | Yes
99 | Usability | Capturing patient data | Failure in heart generation | Rendering mismatch in anatomical model. | Diagnostic support invalid | 3 | 3 | 9 | Tolerable | Implement data validation before rendering. Include error feedback if model generation fails. | 1 | 3 | 3 | Acceptable | No | Yes
100 | Usability | Capturing patient data | Failure in torso generation | Rendering mismatch in anatomical model. | Diagnostic support invalid | 3 | 3 | 9 | Tolerable | Add validation checks for torso mapping. Provide training materials for interpreting projection accuracy. | 1 | 3 | 3 | Acceptable | No | Yes
101 | Usability | Capturing patient data | Failure in detecting electrode position on torso | Electrode markers not identified in VR. | Misleading map or diagnosis | 3 | 3 | 9 | Tolerable | Validate detection logic with known configurations. Show electrode position overlay with real-time feedback. | 1 | 3 | 3 | Acceptable | No | Yes
102 | Usability | Capturing patient data | Measurement capture errors | Torso geometry or electrode position is wrong. | Diagnostic inaccuracy | 3 | 3 | 9 | Tolerable | Implement geometry validation tools. Perform step-by-step acquisition training for users. | 1 | 3 | 3 | Acceptable | No | Yes
103 | User Handling | Device usage | Insufficient GPU processing | Computer lacks resources to run VR smoothly | Lag, missing frames, visual errors | 3 | 3 | 9 | Tolerable | Validate hardware at installation. Include GPU requirements clearly in IFU. | 1 | 3 | 3 | Acceptable | No | Yes
104 | User Handling | Device usage | Cloud case download failure | Poor internet prevents access to patient cases | Delayed diagnosis or loss of access | 3 | 3 | 9 | Tolerable | Buffer cases locally when possible. Notify users and retry downloads. Recommend minimum internet specs. | 1 | 3 | 3 | Acceptable | No | Yes
105 | User Handling | Device usage | Misleading ECG signal in real-time | User interprets raw signal prematurely | Premature or incorrect diagnosis | 3 | 3 | 9 | Tolerable | Block reporting until full signal is processed. | 1 | 3 | 3 | Acceptable | No | Yes
106 | User Handling | Device usage | Unvalidated system modifications | Software is modified without reassessment | Uncontrolled performance deviations | 3 | 3 | 9 | Tolerable | Lock system update and modification controls. Require validation test logs after any system changes. | 1 | 3 | 3 | Acceptable | No | Yes
107 | User Handling | Device usage | Display distortion due to screen size | Image shown with altered proportions | Misinterpretation of heart anatomy | 3 | 3 | 9 | Tolerable | Include scale indicators in VR/2D view. State resolution distortion risks in IFU. | 1 | 3 | 3 | Acceptable | No | Yes
108 | User Handling | Device usage | Lack of VR hardware | System accessed without required VR glasses or compatible PC | Diagnostic mode unavailable | 3 | 3 | 9 | Tolerable | Validate VR hardware at launch. Block access to VR mode without compatible setup. | 1 | 3 | 3 | Acceptable | No | Yes
109 | User Handling | Device usage | Virtual room privacy risk | Room codes shared irresponsibly | Data seen by unauthorized third parties | 3 | 2 | 6 | Tolerable | Require confirmation before sharing room code. Include session participant logs. Warn about sharing risks in UI. | 1 | 2 | 2 | Acceptable | No | Yes
110 | User Handling | Device usage | Real-time pointer delay | Assistants see inaccurate pointer positions | Miscommunication during collaborative diagnosis | 3 | 3 | 9 | Tolerable | Warn users when latency exceeds thresholds. | 1 | 3 | 3 | Acceptable | No | Yes
111 | Execute code | Operating system (Windows 10) | Compatibility errors | Components are incompatible (e.g., drivers, formats, devices). | System failures, inability to visualize or interact with data | 4 | 3 | 12 | Tolerable | Publish minimum hardware specs, validate system compatibility on startup, and offer graceful fallback modes. | 1 | 3 | 3 | Acceptable | No | Yes
112 | Execute code | Operating system (Windows 11) | Compatibility errors | Components are incompatible (e.g., drivers, formats, devices). | System failures, inability to visualize or interact with data | 4 | 3 | 12 | Tolerable | Publish minimum hardware specs, validate system compatibility on startup, and offer graceful fallback modes. | 1 | 3 | 3 | Acceptable | No | Yes

## SOUP Risks (hidden)
Definition of risk | [5 empty] | Risk assessment |  |  |  | Risk acceptance | Risk control (if applicable) |  |  | Risk Control (if applicable) |  |  |  | Risk acceptance | Additional measures |  | Residual risk assessment |  |  | Residual risk control
# | Failure process | Phase | Potential risk | Cause of the risk | Potential harm | Probability | Severity | Detectability | RPN = P x S x D | Acceptable? | Monitoring | Measures to take | Record | Probability | Severity | Detectability | RPN = P x S x D | Acceptable? | Create new risks? | Additional control measures (if applicable) | Residual risk | RPN | Acceptable? | If accepted, information about the patient/user | if not acceptable, risk-benefit assessment
1 | Data processing | Handles dates and times (datetime) | Date/time handling errors | Incorrect date/time formats or calculations | Incorrect or invalid data handling | 1 | 3 | 3 | 9 | No | Unit testing of date/time functions | Implement unit tests; validate date and time inputs and outputs | Test reports | 1 | 2 | 2 | 4 | Yes | No | N/A | N/A | N/A | N/A | N/A | N/A
2 | Data processing | Manages JSON data (json) | Incorrect JSON data handling | Malformed JSON data or content | Data corruption, application crashes | 1 | 3 | 3 | 9 | No | JSON structure validation during development | Validate JSON data structure and content; handle exceptions | Test and exception handling logs | 1 | 2 | 2 | 4 | Yes | No | N/A | N/A | N/A | N/A | N/A | N/A
3 | Data processing | Interacts with OS (os) | OS command errors | Incorrect command execution or paths | System failure, unauthorized access | 1 | 2 | 2 | 4 | Yes | Validation of OS commands and paths during development | Ensure proper permissions and path validation; handle exceptions | Permission validation logs | 1 | 2 | 1 | 2 | Yes | No | N/A | N/A | N/A | N/A | N/A | N/A
4 | Data processing | Accesses interpreter (sys) | Interpreter argument errors | Incorrect arguments or environment variables | Program crashes or unexpected behavior | 1 | 2 | 2 | 4 | Yes | Argument validation during development | Validate arguments and environment variables; handle exceptions | Environment and argument validation logs | 1 | 2 | 1 | 2 | Yes | No | N/A | N/A | N/A | N/A | N/A | N/A
5 | Data processing | Provides platform info (platform) | Platform information errors | Incorrect or inconsistent platform data | Misconfiguration, lack of compatibility | 1 | 2 | 2 | 4 | Yes | Platform info validation across environments | Validate platform information; test across different environments | Platform validation and test reports | 1 | 2 | 1 | 2 | Yes | No | N/A | N/A | N/A | N/A | N/A | N/A
6 | Data processing | Executes system commands (subprocess) | Command execution errors | Invalid or unsafe commands, execution errors | System failures, security vulnerabilities | 1 | 3 | 3 | 9 | No | Command validation and testing during development | Validate commands and handle execution errors; avoid unsafe commands | Execution logs and security checks | 1 | 2 | 2 | 4 | Yes | No | N/A | N/A | N/A | N/A | N/A | N/A
7 | Data processing | Provides time functions (time) | Time-related errors | Incorrect time calculations or conversions | Incorrect timing, data corruption | 1 | 2 | 2 | 4 | Yes | Time function validation during development | Validate time usage and handle exceptions | Time function validation reports | 1 | 2 | 1 | 2 | Yes | No | N/A | N/A | N/A | N/A | N/A | N/A
8 | Data processing | Math/stat calculations (numpy) | Calculation errors | Incorrect calculations or dependencies | Incorrect data analysis, misleading results | 1 | 3 | 3 | 9 | No | Extensive unit testing and dependency review | Conduct extensive unit testing; validate results; review dependencies | Unit test and dependency review reports | 1 | 2 | 2 | 4 | Yes | No | N/A | N/A | N/A | N/A | N/A | N/A
9 | Data processing | Data analysis (pandas) | Data manipulation errors | Incorrect data handling or dependencies | Data corruption, incorrect analysis | 1 | 3 | 3 | 9 | No | Data validation and unit testing during development | Validate data and conduct unit tests; review dependencies | Data validation and unit test reports | 1 | 2 | 2 | 4 | Yes | No | N/A | N/A | N/A | N/A | N/A | N/A
10 | Data processing | Sends HTTP requests (request) | HTTP request errors | Incorrect responses or handling of HTTP errors | Communication failures, data loss | 1 | 3 | 3 | 9 | No | HTTP response validation and error handling during development | Validate HTTP responses; handle errors and exceptions | HTTP response logs and error handling reports | 1 | 2 | 2 | 4 | Yes | No | N/A | N/A | N/A | N/A | N/A | N/A
11 | Data processing | Works with 3D meshes (trimesh) | 3D mesh handling errors | Incorrect 3D data input or processing | Visualization errors, data loss | 1 | 2 | 2 | 4 | Yes | 3D data validation during development | Validate 3D data input and output; conduct integrity tests | 3D data validation and integrity test reports | 1 | 2 | 1 | 2 | Yes | No | N/A | N/A | N/A | N/A | N/A | N/A
12 | Data processing | Web framework (flask) | Web application errors | Security vulnerabilities or functional issues | Data breaches, application failure | 1 | 3 | 3 | 9 | No | Security and functionality testing during development | Ensure web application security; perform functionality and load tests | Security and functionality test reports | 1 | 2 | 2 | 4 | Yes | No | N/A | N/A | N/A | N/A | N/A | N/A
13 | Execute code | Operating system (Windos 10) | Compatibility errors | Software not compatible with OS version | Software failures, security vulnerabilities | 2 | 3 | 3 | 18 | No | Compatibility testing during development and updates | Ensure compatibility with software; keep security updates current | Compatibility and update reports | 1 | 2 | 2 | 4 | Yes | No | N/A | N/A | N/A | N/A | N/A | N/A
14 | Execute code | Operating system (Windos 11) | Compatibility errors | Software not compatible with OS version | Software failures, security vulnerabilities | 2 | 3 | 3 | 18 | No | Compatibility testing during development and updates | Ensure compatibility with software; keep security updates current | Compatibility and update reports | 1 | 2 | 2 | 4 | Yes | No | N/A | N/A | N/A | N/A | N/A | N/A
Note on B3: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: It seems to me that the names listed are of the names of the SOUPs themselves and not the process in which the fault can occur, please adapt considering the examples my colleague Celeste emailed. For example, the process in which the failure occurs may be the integration of SOUP?

## Severity score (hidden)
Common terms | Possible description
Catastrophic / Fatal | Results in death
Critical | Results in permanent impairment or irreversible injury
Serious / Major | Results in injury or impairment requiring medical or surgical intervention
Minor | Results in temporary injury or impairment not requiring medical or surgical intervention
Negligible | Results in inconvenience or temporary discomfort
Note on C1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: I find it unnecessary to put severity score here given that the same is found in the risk management plan. Still, if it's Spika Tech's option to put, then the scores for probability and detectability should be placed also.

## Risk Evaluation (hidden)
risk | Implementation | Need action?
Low | n/a | Forks
Medium | waiting | No
High | Ongoing
n/a | Concluded

## Risk Manag. Team
Function | Responsibilities | Represented by
Executive Director | Assess potential risks from the perspective of usage and customer complaints, such as feedback. | Fernando Pozo
Quality guarantee | Evaluate possible risks from the point of view of the clinical area and compliance with the requirements defined in the regulations and laws in force for the product. | Iván Pérez
IT & Robotics | Assess potential risks from a hardware implementation perspective. Evaluate potential risks from a software implementation perspective. Assess possible risks from a production point of view and “control procedures for non-conforming products” and implementation of the “corrective and preventive action control strategy” in the production process. | José Martín
Investigation and development | Assess potential risks from a software implementation perspective. Assist in the creation, review and approval of risk management activities. | Vicente Gallego
Sales | Responsible for providing information, as needed, regarding marketing, usage, customer feedback, and new risks. | Raúl Cordón
Healthcare professionals | Assess potential risks from the perspective of the clinical environment and clinical practices and provide this information to the manufacturer. | José Ángel Cabrera
PRRC | Coordinate the Team for the Identification and Assessment of Potential Hazards. | Celeste Oliveira

## Criteria
Risk Index (RI)
Severity of the harm | Description | Score | Risk assessment - Level of risk
Negligible | Results in inconvenience or temporary discomfort | 1
Minor | Results in temporary injury or impairment not requiring medical or surgical intervention | 2 | Value (Severity x Occurence) | Acceptance
Serious / Major | Results in injury or impairment requiring medical or surgical intervention | 3 | 1 | Acceptable
Critical | Results in permanent impairment or irreversible injury | 4 | 2 | Acceptable
Catastrophic / Fatal | Results in death | 5 | 3 | Acceptable
|  |  | 4 | Acceptable
Probability of the harm | Description | Score | 5 | Acceptable
Improbable | <10−6 | 1 | 6 | Tolerable
Remote | <10−5 and ≥10−6 | 2 | 8 | Tolerable
Occasional | <10−4 and ≥10−5 | 3 | 9 | Tolerable
Probable | <10−3 and ≥10−4 | 4 | 10 | Tolerable
Frequent | ≥10−3 | 5 | 12 | Tolerable
|  |  | 15 | Unnacceptable
Residual Risk (RR) |  |  | 16 | Unnacceptable
Severity of the harm |  | Score | 20 | Unnacceptable
Negligible |  | 1 | 25 | Unnacceptable
Minor |  | 2
Serious / Major |  | 3
Critical |  | 4
Catastrophic / Fatal |  | 5
Occurrence probability |  | Score | Acceptable Residual Risk Scores | Acceptance conditions
Improbable |  | 1 | RR ≤ 5 | N/A
Remote |  | 2 | 6 ≤ RR < 15 | Alert in IFU
Occasional |  | 3 | RR ≥ 15 | Benefit > Risk Alert in IFU
Probable |  | 4
Frequent |  | 5