Annex 2_VERSION_05

- CAPA Report

| LAYER REPORT | CODE |
| --- | --- |
|  | AC-003/25 |

| Person opening the action: | Iván Pérez |  |  |  |  |  |  | Department |  |  |  | Quality Assurance |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Detection date | 15/07/2025 |  |  |  |  |  |  | Origin |  | Notified Body Audit |  |  |
| Description of the finding | The methodology used in the FMEA report for calculating the Risk Priority Number (RPN) is based on the multiplication of severity, probability, and detectability and does not align with the Risk Management procedure PNT-002.04, which defines the RPN calculation as the product of severity and probability only |  |  |  |  |  |  |  |  |  |  |  |
| Root cause analysis | A discrepancy has been identified between the methodology applied in the FMEA report and the risk management procedure PNT-002.04. While the procedure defines the calculation of the RPN as the multiplication of severity and probability (RPN = S × P), the FMEA report uses the traditional formula S × P × D (detectability). This inconsistency is due to the lack of alignment between the template used for the FMEA and the internal criteria defined in the procedure. |  |  |  |  |  |  |  |  |  |  |  |
| Immediate correction | Record the non-conformity in the quality system and provide traceability of the follow-up. |  |  |  |  |  |  |  |  |  |  |  |
| Corrective/preventive action | • Updating the standard FMEA analysis template to correctly reflect the internal methodology for calculating the RPN (S × P). • Review of previous risk analyses to identify possible inconsistent documents. |  |  |  |  |  |  |  |  |  |  |  |
| Verification of the impact of the action on the product |  |  |  |  |  |  | Yes (if you check this box, explain below) |  |  |  | No |  |
|  | Is there an impact on regulatory requirements? |  |  |  |  |  | ☐ |  |  |  | ☒ |  |
|  | Is there an impact on safety and operation? |  |  |  |  |  | ☐ |  |  |  | ☒ |  |
|  | Impact assessment: |  |  |  |  |  | N/A |  |  |  |  |  |
| Responsible for the action | Iván Pérez |  |  |  |  |  |  |  | Department |  |  | Quality Assurance |
| Deadline | 31/10/2025 |  |  |  |  |  |  |  |  |  |  |  |
| To be completed by the person who opens | Signature and date |  |  |  |  |  |  |  |  |  |  |  |
| To be completed by the quality manager: | Signature and date Date: |  |  |  |  |  |  |  |  |  |  |  |
| Evaluation of the effectiveness of the change – To be completed when the change is made |  |  |  |  |  |  |  |  |  |  |  |  |
| Has the change been implemented successfully? |  | ☐ |  | Yes |  |  |  |  |  |  |  |  |
|  |  | ☐ |  | No | Observations |  |  |  |  |  |  |  |
| Effectiveness of change – To be completed when the change is made |  |  |  |  |  |  |  |  |  |  |  |  |
| Closing carried out by: |  |  |  |  |  | Date |  |  |  |  |  |  |